# Arkeo Multichannel API

v0.4.1
20-01-2026

Contents

Introduction	2

Summary of available commands	2

List of commands	3

Command: SetActiveChannel	3

Command: GetActiveChannel	3

Command: SetChannelSettings	3

Command: GetChannelSettings	3

Command: StartChannel	3

Command: StopChannel	4

Command: ForceJV	4

Command: GetChannelState	4

Command: GetLatestJV	4

Command: GetIV	5

Command: GetSensorsInfo	5

Command: GetSensors	5

The JSON settings string	6

The JSON state string	9

The Sensor JSON String	10

Examples	11

Python	11

Changelog	12

# Introduction

This API facilitates remote control of the Arkeo Multichannel using TCP on port 6340, enabling users to perform a variety of commands manage measurements.

A command must sent as a JSON string in the form
{'command': command, 'parameter': parameter}

Each command must be preceded by a 4-byte integer indicating the total length of the command. After each command, a reply is always sent back. Its format depends on the command (see list below). If an unknown command is sent, the string “Not a valid command” is send instead.

# Summary of available commands

| Command | Description |
| --- | --- |
| SetActiveChannel | Set the channel upon which all subsequent actions are performed |
| GetActiveChannel | Get the active channel |
| SetChannelSettings | Modify the settings |
| GetChannelSettings | Retrieve the settings |
| StartChannel | Start measurement |
| StopChannel | Stop the measurement process |
| ForceJV | Force a JV measurement |
| GetChannelState | Retrieve the current state |
| GetLatestJV | Get the voltage and current of last performed JV |
| GetIV | Get the live voltage and current of all channels |
| GetSensors | Get the live sensor values of all connected sensors |

## List of commands

### Command: SetActiveChannel

Description: Set the channel upon which all succeeding actions are performed.
Parameters:

- channel_id (integer): The ID of the channel.

Response:

- channel_id (integer): The ID of the channel.

### Command: GetActiveChannel

Description: Get the active channel.
Parameters:

Response:

- channel_id (integer): The ID of the channel.

### Command: SetChannelSettings

Description: Set the JSON configuration string for active channel.
Parameters:

- settings (JSON string): The configuration settings in JSON format (see chapter: The JSON settings string).

Response:

- status (string): OK

### Command: GetChannelSettings

Description: Retrieve the JSON configuration string of active channel.
Parameters:

Response:

- settings (JSON string): The configuration settings in JSON format (see chapter: The JSON settings string).

### Command: StartChannel

Description: Start the measurement process for active channel.
Parameters:

Response:

- status (string): OK

### Command: StopChannel

Description: Stop the measurement process for active channel.
Parameters:

Response:

- status (string): OK

### Command: ForceJV

Description: Force a JV measurement on active channel if channel is in tracking mode.
Parameters:

- channel_id (integer): The ID of the active channel.

Response:

- status (string): OK

### Command: GetChannelState

Description: Retrieve the current state of a specific channel (e.g., running, stopped, JV, tracking).
Parameters:

Response:

- state (string): JSON of the state of the active channel (see chapter: The JSON state string)

### Command: GetLatestJV

Description: Retrieve the voltage and current values of the latest JV.
Parameters:

Response:

- JV Data (string):
v_fw1|j_fw 1|...|v_fw n|j_fw n|| v_rv1|j_rv 1|...|v_rv n|j_rv n

Note:

Forward and Reverse scans are separated by || (double pipe character)
Voltage and Current arrays are interleaved and separated by | (single pipe character)

Voltage unit: V

Current unit A/cm²

### Command: GetIV

Description: Retrieve the live voltage and current values of all channels.
Parameters:

Response:

- Voltage and Current values (string):
v1|j1|...|vn|jn

Note:

Voltage and Current arrays are interleaved and separated by | (single pipe character)

Voltage unit: V

Current unit A/cm²

### Command: GetSensorsInfo

Description: Retrieve sensors configuration.
Parameters:

Response:

- Sensor values (string):
JSON string containing: Channel, Name, Type, Unit, sensor specific settings

Note:

See The Sensor JSON String

### Command: GetSensors

Description: Retrieve the live values of all sensors.
Parameters:

Response:

- Sensor values (string):
s1|...|sn|

Note:

Sensor values are separated by | (single pipe character)

Use Command: GetSensorsInfo to associate the values to the correct sensor and unit.

### Command: GetEvironments

Description: Retrieve the all environments and the assosciated sensors.
Parameters:

Response:

[

{

"name":"Env 1",

"sensors_id":[

"Humidity;1;Temperature",

"O2;7;Oxygen",

"Temperature;0;Humidity"

]

}

]

Note:

The sensors_id list contains unique IDs for each sensor. Use the ‘GetSensorInfo’ command to assosciated the ID to the appropriate sensor.

### Command: SetChannelEnvironment

Description: Sets the environment of the active channel.
Parameters:

- environment (string): The name of the environment. Set an empty string to not link the channel to any environment

Response:

- OK

Note: Use the GetEvironments command to get a list of valid environment names. Setting a name which is not in this list returns an error and does not update the channel.

## The JSON settings string

The settings of a channel are represented in a JSON string. When using the Command: SetChannelSettings, this is value that the software expects. It is recommended to use the Command: GetChannelSettings to read the actual settings and modify those settings.

Example

Below is an example of a channel settings string. A detailed description can be found on the next page.

{

"Index":"1A",

"Enable":true,

"User":"Cicci Research",

"Device":"Si cell",

"Channel":{

"VoltageLimit":"10 V",

"CurrentLimit":0,

"InvertedStructure":false

},

"JV":{

"Vmin (V)":-0.1,

"Vmax (V)":1.2,

"Step (mV)":20,

"ScanRate (mV/s)":100,

"VocDetect":true,

"Overvoltage (%)":0,

"ScanOrder":"FW then RV"

},

"Tracking":{

"TrackEnable":true,

"Algorithm":"MPPT",

"Perturbation (V)":0.02,

"ConstantOutput":0,

"SaveInterval (s)":10,

"jvInterval":{"Value":10,"Unit":"min"},

"TestDuration":{"Value":100,"Unit":"hours"}

},

"Cell":{

"Type":"Cell",

"Area (cm2)":1,

"NrCells":1,

"NrW cells":1,

"W-cellArea (cm2)":1

},

"Note":""

}

Description

| Parameter | Description | Example | Unit | Type |
| --- | --- | --- | --- | --- |
| Index | Unique channel identifier | 0 |  | integer |
| Enable | Enables or disables the channel | true |  | boolean |
| User | Name of the user | User |  | string |
| Device | Device Name | Sample |  | string |
| Channel |  |  |  | object |
| - VoltageLimit1 | Maximum Voltage | 10 V |  | enum |
| - CurrentLimit | Current Range | 0 |  | integer |
| - InvertedStructure | Inverts the applied voltage | false |  | boolean |
| JV |  |  |  | object |
| - Vmin (V) | Minimum voltage | -0.1 | V | float |
| - Vmax (V) | Maximum voltage | 2 | V | float |
| - Step | Voltage increment per step | 20 | mV | integer |
| - ScanRate (mV/s) | Rate at which the voltage is applied | 100 | mV/s | float |
| - VocDetect | Enables the detection of open-circuit voltage | true |  | boolean |
| - ScanOrder2 | Order of scanning | FW then RV |  | enum |
| Tracking |  |  |  | object |
| - TrackEnable | Enables tracking | true |  | boolean |
| - Algorithm3 | Specifies the algorithm used for tracking | MPPT |  | enum |
| - Perturbation (V) | Voltage differential for the tracking algorithm | 0.01 | V | float |
| - ConstantOutput | Setting to maintain a constant output | 0.2 | * | float |
| - SaveInterval (s) | Time between saved data poitns | 10 | s | integer |
| - jvInterval.Value | Time between JV scans | 10 |  | float |
| - jvInterval.Unit4 | Unit for JV interval | min |  | enum |
| - TestDuration.Value | Duration of the tracking test | 100 |  | float |
| - TestDuration.Unit4 | Unit for Test Duration | hours |  | enum |
| Cell |  |  |  | object |
| - Type5 | Cell Type | Cell |  | enum |
| - Area (cm2) | Area of the cell | 1 | cm² | float |
| - NrCells | Number of cells | 1 |  | integer |
| - NrW cells | Number of W cells | 1 |  | integer |
| - W-cellArea (cm2) | The area of each W cell | 1 | cm² | float |
| Note | Free to use field for comments |  |  | string |

Notes for settings JSON

Detailed explanation and enums of the settings JSON

#### Voltage Limit

The voltage limit of each board is 10 V. 20 V can be reached by connecting 2 boards in parallel. Connect the positive contacts of 2 boards to your device and the negative contacts to each other.

#### Scan Order enum

0: FW then RV

1: RV then FW

2: Forward Only

3: Reverse Only

#### Tracking Algorithm enum

0: Open circuit

1: Short circuit

2: MPPT

3: MPPT-Stab

4: MPPT INC

5: Fixed Voltage

6: Fixed Voltage (no track)

7: Fixed Current

8: JV

#### Time unit enum

0: sec

1: min

2: hours

#### Cell Type enum

0: Cell

1: Parallel Module

2: Z Module

3: W Module

## The JSON state string

The state of a channel is represented in a JSON string. This string is read-only and is obtained using the Command: GetChannelState.

Example

Below is an example of a state string. A detailed description can be found below.

{"Enable":false,

"Channel":"1A",

"User":"User",

"Measurement":"JV",

"Direction":"Forward",

"State":"Running",

}

Description

| Parameter | Description | Example | Data Type |
| --- | --- | --- | --- |
| Enable | Enables or disables the channel | true | Boolean |
| Channel | ID of the channel | 1A | String |
| User | Name of the user | User | String |
| Measurement | State of the Measurement | JV | String |
| Direction | Current direction of the JV | Forward | String |
| State | Current state of the channel - Idle - Ready to start - Running - Paused - Stopped - Error | Running | String |

## The Sensor JSON String

Sensor data is retreived as a JSON string containing an array of all configured sensors. The specific section contains sensor specific configuration and/or calibration data. It can be empty.

Example

[

{

"type":"Pyranometer",

"common":{

"Channel":2,

"Type":"Luminosity",

"Unit":"W/m2",

"Custom Name":"Pyranometer (ai2)"

},

"specific":{

"Calibration (V/(W/m2))":9.13E-6,

"Gain":200

}

},

{

"type":"SHT3x-ARP Temperature",

"common":{

"Channel":1,

"Type":"Temperature",

"Unit":"degC",

"Custom Name":""

},

"specific":{

}

}

]

# Examples

## Python

import socket

import json

def exampleCommand(command, data=''):

with socket.socket(socket.AF_INET, socket.SOCK_STREAM) as sock:

sock.connect(('localhost', 6340))

message = json.dumps({'command': command, 'data': data})

# Prepare the length of the string

length = len(message)

length_bytes = length.to_bytes(4, byteorder='big')  # 4 bytes to represent the length

sock.sendall(length_bytes) # Send the length of the string

sock.sendall(message.encode('utf-8')) # Send the string

# Receive response

response = sock.recv(1024)

print('Received:', response.decode('utf-8'))

exampleCommand('GetChannelSettings')

# Changelog

v0.1
Initial version

v0.2
Added ForceJV and GetChannelState

v0.3
Added GetLatestJV, GetIV and GetSensors

v0.4
Reorganized the settings JSON and clarified the documentation
Added JSON schema validation

v0.4.1
Now allows multiple connections in parallel
Added GetSensorsInfo command
GetSensors now sends physical values instead of raw voltages
GetSensors no longer sends all sensor channels, but only configured channels
Index field in the settings JSON changed from string to integer

v0.4.2
Added GetEnvironments
Added SetChannelEnvironment

v0.4.3
Added Day-Night object to json settings
Added GetGlobalDayNight and SetGlobalDayNight