# Arkeo Multichannel API

v0.4
21-08-2025

Contents

Introduction	2

Summary of available commands	2

List of commands	3

Command: SetActiveChannel	3

Command: GetActiveChannel	3

Command: SetChannelSettings	3

Command: GetChannelSettings	3

Command: StartChannel	3

Command: StopChannel	4

Command: ForceJV	4

Command: GetChannelState	4

Command: GetLatestJV	4

Command: GetIV	5

Command: GetSensors	5

The JSON settings string	6

Example	6

Description	7

Notes for settings JSON	8

The JSON state string	9

Example	9

Description	9

Examples	10

Python	10

Changelog	11

# Introduction

This API facilitates remote control of the Arkeo Multichannel using TCP on port 6340, enabling users to perform a variety of commands manage measurements.

A command must sent as a JSON string in the form
{'command': command, 'parameter': parameter}

Each command must be preceded by a 4-byte integer indicating the total length of the command. After each command, a reply is always sent back. Its format depends on the command (see list below). If an unknown command is sent, the string “Not a valid command” is send instead.

# Summary of available commands

| Command | Description |
| --- | --- |
| SetActiveChannel | Set the channel upon which all subsequent actions are performed |
| GetActiveChannel | Get the active channel |
| SetChannelSettings | Modify the settings |
| GetChannelSettings | Retrieve the settings |
| StartChannel | Start measurement |
| StopChannel | Stop the measurement process |
| ForceJV | Force a JV measurement |
| GetChannelState | Retrieve the current state |
| GetLatestJV | Get the voltage and current of last performed JV |
| GetIV | Get the live voltage and current of all channels |
| GetSensors | Get the live sensor values of all connected sensors |

## List of commands

### Command: SetActiveChannel

Description: Set the channel upon which all succeeding actions are performed.
Parameters:

- channel_id (integer): The ID of the channel.

Response:

- channel_id (integer): The ID of the channel.

### Command: GetActiveChannel

Description: Get the active channel.
Parameters:

Response:

- channel_id (integer): The ID of the channel.

### Command: SetChannelSettings

Description: Set the JSON configuration string for active channel.
Parameters:

- settings (JSON string): The configuration settings in JSON format (see chapter: The JSON settings string).

Response:

- status (string): OK

### Command: GetChannelSettings

Description: Retrieve the JSON configuration string of active channel.
Parameters:

Response:

- settings (JSON string): The configuration settings in JSON format (see chapter: The JSON settings string).

### Command: StartChannel

Description: Start the measurement process for active channel.
Parameters:

Response:

- status (string): OK

### Command: StopChannel

Description: Stop the measurement process for active channel.
Parameters:

Response:

- status (string): OK

### Command: ForceJV

Description: Force a JV measurement on active channel if channel is in tracking mode.
Parameters:

- channel_id (integer): The ID of the active channel.

Response:

- status (string): OK

### Command: GetChannelState

Description: Retrieve the current state of a specific channel (e.g., running, stopped, JV, tracking).
Parameters:

Response:

- state (string): JSON of the state of the active channel (see chapter: The JSON state string)

### Command: GetLatestJV

Description: Retrieve the voltage and current values of the latest JV.
Parameters:

Response:

- JV Data (string):
v_fw1|j_fw 1|...|v_fw n|j_fw n|| v_rv1|j_rv 1|...|v_rv n|j_rv n

Note:

Forward and Reverse scans are separated by || (double pipe character)
Voltage and Current arrays are interleaved and separated by | (single pipe character)

Voltage unit: V

Current unit A/cm²

### Command: GetIV

Description: Retrieve the live voltage and current values of all channels.
Parameters:

Response:

- Voltage and Current values (string):
v1|j1|...|vn|jn

Note:

Voltage and Current arrays are interleaved and separated by | (single pipe character)

Voltage unit: V

Current unit A/cm²

### Command: GetSensors

Description: Retrieve the live voltags of all sensors.
Parameters:

Response:

- Sensor values (string):
v1|...|vn|

Note:

Sensor voltages are separated by | (single pipe character)

Voltage unit: V

## The JSON settings string

The settings of a channel are represented in a JSON string. When using the Command: SetChannelSettings, this is value that the software expects. It is recommended to use the Command: GetChannelSettings to read the actual settings and modify those settings.

### Example

Below is an example of a channel settings string. A detailed description can be found on the next page.

{

"Index":"1A",

"Enable":true,

"User":"Cicci Research",

"Device":"Si cell",

"Channel":{

"VoltageLimit":"10 V",

"CurrentLimit":0,

"InvertedStructure":false

},

"JV":{

"Vmin (V)":-0.1,

"Vmax (V)":1.2,

"Step (mV)":20,

"ScanRate (mV/s)":100,

"VocDetect":true,

"Overvoltage (%)":0,

"ScanOrder":"FW then RV"

},

"Tracking":{

"TrackEnable":true,

"Algorithm":"MPPT",

"Perturbation (V)":0.02,

"ConstantOutput":0,

"SaveInterval (s)":10,

"jvInterval":{"Value":10,"Unit":"min"},

"TestDuration":{"Value":100,"Unit":"hours"}

},

"Cell":{

"Type":"Cell",

"Area (cm2)":1,

"NrCells":1,

"NrW cells":1,

"W-cellArea (cm2)":1

},

"Note":""

}

### Description

| Parameter | Description | Example | Unit | Type |
| --- | --- | --- | --- | --- |
| Index | Unique channel identifier | 1A |  | string |
| Enable | Enables or disables the channel | true |  | boolean |
| User | Name of the user | User |  | string |
| Device | Device Name | Sample |  | string |
| Channel |  |  |  | object |
| - VoltageLimit1 | Maximum Voltage | 10 V |  | enum |
| - CurrentLimit | Current Range | 0 |  | integer |
| - InvertedStructure | Inverts the applied voltage | false |  | boolean |
| JV |  |  |  | object |
| - Vmin (V) | Minimum voltage | -0.1 | V | float |
| - Vmax (V) | Maximum voltage | 2 | V | float |
| - Step | Voltage increment per step | 20 | mV | integer |
| - ScanRate (mV/s) | Rate at which the voltage is applied | 100 | mV/s | float |
| - VocDetect | Enables the detection of open-circuit voltage | true |  | boolean |
| - ScanOrder2 | Order of scanning | FW then RV |  | enum |
| Tracking |  |  |  | object |
| - TrackEnable | Enables tracking | true |  | boolean |
| - Algorithm3 | Specifies the algorithm used for tracking | MPPT |  | enum |
| - Perturbation (V) | Voltage differential for the tracking algorithm | 0.01 | V | float |
| - ConstantOutput | Setting to maintain a constant output | 0.2 | * | float |
| - SaveInterval (s) | Time between saved data poitns | 10 | s | integer |
| - jvInterval.Value | Time between JV scans | 10 |  | float |
| - jvInterval.Unit4 | Unit for JV interval | min |  | enum |
| - TestDuration.Value | Duration of the tracking test | 100 |  | float |
| - TestDuration.Unit4 | Unit for Test Duration | 1 |  | enum |
| Cell |  |  |  | object |
| - Type5 | Cell Type | Cell |  | enum |
| - Area (cm2) | Area of the cell | 1 | cm² | float |
| - NrCells | Number of cells | 1 |  | integer |
| - NrW cells | Number of W cells | 1 |  | integer |
| - W-cellArea (cm2) | The area of each W cell | 1 | cm² | float |
| Note | Free to use field for comments |  |  | string |

### Notes for settings JSON

Detailed explanation and enums of the settings JSON

#### Voltage Limit

The voltage limit of each board is 10 V. 20 V can be reached by connecting 2 boards in parallel. Connect the positive contacts of 2 boards to your device and the negative contacts to each other.

#### Scan Order enum

0: FW then RV

1: RV then FW

2: Forward Only

3: Reverse Only

#### Tracking Algorithm enum

0: Open circuit

1: Short circuit

2: MPPT

3: MPPT-Stab

4: MPPT INC

5: Fixed Voltage

6: Fixed Voltage (no track)

7: Fixed Current

8: JV

#### Time unit enum

0: seconds

1: minutes

2: hours

#### Cell Type enum

0: Cell

1: Parallel Module

2: Z Module

3: W Module

## The JSON state string

The state of a channel is represented in a JSON string. This string is read-only and is obtained using the Command: GetChannelState.

### Example

Below is an example of a state string. A detailed description can be found below.

{"Enable":false,

"Channel":"1A",

"User":"User",

"Measurement":"JV",

"Direction":"Forward",

"State":"Running",

}

### Description

| Parameter | Description | Example | Data Type |
| --- | --- | --- | --- |
| Enable | Enables or disables the channel | true | Boolean |
| Channel | ID of the channel | 1A | String |
| User | Name of the user | User | String |
| Measurement | State of the Measurement | JV | String |
| Direction | Current direction of the JV | Forward | String |
| State | Current state of the channel - Idle - Ready to start - Running - Paused - Stopped - Error | Running | String |

# Examples

## Python

import socket

import json

def exampleCommand(command, data=''):

with socket.socket(socket.AF_INET, socket.SOCK_STREAM) as sock:

sock.connect(('localhost', 6340))

message = json.dumps({'command': command, 'data': data})

# Prepare the length of the string

length = len(message)

length_bytes = length.to_bytes(4, byteorder='big')  # 4 bytes to represent the length

sock.sendall(length_bytes) # Send the length of the string

sock.sendall(message.encode('utf-8')) # Send the string

# Receive response

response = sock.recv(1024)

print('Received:', response.decode('utf-8'))

exampleCommand('GetChannelSettings')

# Changelog

v0.1
Initial version

v0.2
Added ForceJV and GetChannelState

v0.3
Added GetLatestJV, GetIV and GetSensors

v0.4
Reorganized the settings JSON and clarified the documentation
Added JSON schema validation